Standard Test Method for

Rotational Viscosity of  Heavy Duty Diesel Drain Oils at 100( C

This standard is issued under the fixed designation X XXXX; the number immediately following the designation indicates the year of original adoption or, in the case of revision, the year of last revision. A number in parentheses indicates the year of last reapproval. A superscript epsilon () indicates an editorial change since the last revision or reapproval.

- Scope

  - This test method covers the determination of the rotational viscosity and the shear thinning properties of heavy duty diesel engine drain oils at 100C, in the shear rate range of approximately 10 – 300 s-1 , in the shear stress range of approximately 0.1-10 Pa and the viscosity range of approximately 12-35 mPa-s.  The shear thinning property of a diesel drain oil is a measure of the oils ability to handle soot as evidenced by an increase in viscosity as determined in ASTM D5967-99a.  Rotational viscosity values can be compared at identical shear rates by this method (references 1,2)

  - This standard does not purport to address all of the safety concerns, if any, associated with its use. It is the responsibility of the user of this standard to establish appropriate safety and health practices and to determine the applicability of regulatory limitations prior to use.

- Referenced Documents

  - ASTM Standards:

ASTM D5967 - 99a Standard Test Method for Evaluation of Diesel Engine Oils in T-8 Diesel Engine

ASTM D6299 - 00 Standard Practice for Applying Statistical Quality Assurance Techniques to Evaluate Analytical Measurement System Performance

ASTM D4057-95(2000) Standard Practice for Manual Sampling of Petroleum and Petroleum Products

- Terminology

  - Terms Specific to this Standard:

Rotational Viscosity – the viscosity obtained by use of this method

Rate Index -The exponent, c, in the expression relating shear rate and shear stress:

Shear Stress = b(shear rate)c                                                       (1)

or                    ln(shear stress) = ln b + c ln(shear rate)                                             (2)

A rate index of c=1 signifies  Newtonian fluid behavior.  Values less than one indicate increasing non-Newtonian , shear thinning behavior (reference 2).

VIS100 INC – rotational viscosity at shear rate of 100 s-1, increasing shear stress or shear rate sweep

VIS100 DEC - rotational viscosity at shear rate of 100 s-1, decreasing shear stress or shear rate sweep

- Summary of Test Method

  - The sample is placed in a controlled stress or controlled shear  rate rheometer/viscometer at 100C.  The sample is presheared at 10 s-1 for 30 s followed by preheating at 100C for 10 minutes. An increasing shear rate (approximately 10-300s-1) or shear stress (.1-10 Pa) sweep is run followed by a decreasing sweep.  The rotational viscosity for each step (increasing and decreasing) at 100 s-1 shear rate is interpolated from the viscosity versus shear rate data table.  The rate index, as a measure of shear thinning, is calculated from a plot of ln (shear stress) versus ln (shear rate).

- Significance and Use

Rotational viscosity measurements allow the determination of the non-Newtonian, shear thinning property of drain oils.  This property is a result of the ability of the oil to handle soot as evaluated in the ASTM D5967-99a engine test method.  Oil thickening is a consequence of the oils inability to handle soot.   Rotational viscosity values can be compared at  identical shear rates by this method  (References1,2).

- Apparatus

This test method uses rheometers/viscometers of the controlled stress or controlled rate mode of operation.  The method requires the use of cone and plate or concentric cylinders measuring geometry. capable of operating in the range of approximately 0.1-10 Pa for shear stress and 10-300 s-1 for shear rate.  Instrument data logging or software shall be capable of delivering shear stress versus shear rate data and viscosity versus shear rate data in tabular form.  Temperature shall be controlled to 100C  0.2C at equilibrium.  Some rheometers have a 99.9C set point limit and would be acceptable for this method.

- Reagents and Materials

  - Standard Oil, calibrated in viscosity in the range of 12 – 35  mPa-s at 100C.

- Sampling, Test Specimens, and Test Units

  - Ensure the test specimen is homogeneous.  Engine sampling is generally specified in the test method, for example ASTM D5697-99a.  Manual sampling from the container can be done following ASTM D 4057.

- Preparation of Apparatus

  - Prepare the apparatus in accordance with manufacturers directions. The apparatus shall be capable of viscosity measurement to within 5% for a standard oil and a rate index value of 0.98-1.02 indicating a Newtonian fluid for the standard oil.

- Calibration and Standardization

  - A Newtonian viscosity standard in the range 12 to 35 mPa-s at 100C shall be used to verify instrument calibration.  Run the procedure as in section 12  A plot of shear stress (Pa) vs. shear rate (s-1) shall be linearly regressed to yield a slope and intercept.

Results should be:

intercept ,  <0.1 Pa

slope = viscosity value within 5% of certified value, mPa-s

correlation coefficient , r2  > .9998

The instrument manufacturer should be contacted if the instrument is out of calibration.

  - New SAE 15W-40 oil shall be used as a daily control chart standard.  Run the procedure (section 12) and perform an analysis (section 13).  Results shall be a rate index between .98-1.02 and a viscosity value at 100 s-1, mPa-s, VIS100.  Control chart the values of VIS100 and rate index.  The procedure shall be checked and the instrument calibration rechecked if the reference oil does not fall within control limits. ASTM D6299 shall be used as a guide in this area

  - Some instruments and geometries will exhibit significant noise at low shear stress and /or shear rate levels.  This may be determined by plotting viscosity vs. shear rate or viscosity vs. shear stress for measurements of the standard oil.  A horizontal line is obtained in regions far from noise.  Appendix 1 shows two examples of this type of plot.  The minimum shear rate or shear stress to use in the analysis of data can then be determined for the particular instrument and geometry.

- Conditioning

  - Shake all new and used oil samples using the following procedure.  Do not prepare more than two samples at one time for one instrument.

  - Ensure cap is tight on container.

  - Shake vigorously by hand for 30 s.  Wait 60 s for air bubbles to dissipate.

  - A specimen of the sample shall be taken for analysis promptly following the shaking and dissipation procedure.

- Procedure

Run the procedure per the instrument geometry requirements and the manufacturer recommendations to obtain shear stress versus shear rate data in the appropriate ranges of 0.1-10 Pa and 10-300 s-1.  The order of steps is as follows:

  - Load Sample

  - Equilibrate at 100C (maximum 10 minutes)

  - Preshear sample at 10 s-1 for 30 s

  - Stop preshear

  - Preheat sample at 100C for 10 minutes

  - Run increasing stress or rate sweep for duration of approximately 10 minutes to generate data of shear stress, shear rate and viscosity followed immediately by the next step.  The run time will vary somewhat among different instruments and procedures.  Times as low as 2 minutes and as high as 20 minutes have been utilized successfully to run this method.

  - Run decreasing stress or rate sweep for duration of approximately 10 minutes to generate data of shear stress, shear rate and viscosity.

  - Clean sample from instrument per manufacturers instructions.  Cone and plate systems may be cleaned by wiping with a rag followed by a solvent clean such as textile spirits.

- Calculation or Interpretation of Results                                                                      Analyze the increasing and decreasing sweeps separately

  - Import the shear stress (Pa), shear rate (s-1) and viscosity (mPa-sec) into a spreadsheet program.  This calculation may be done with a calculator.  Delete data below the noise limit as determined in section 10.3

  - Calculate two additional columns for ln (shear rate) and ln (shear stress).

  - Plot the ln stress vs ln rate columns as a scatter plot with ln rate on the x-axis.

  - Fit a least squares linear regression to the data plot – obtain the equation of the line and correlation coefficient.  This line follows equation (2).

  - Obtain the slope of the line, c which is the rate index, to three decimal places. the correlation coefficient r2 to four decimal places and the intercept as  ln b to four decimal places.

  - Calculate the viscosity at 100 s-1 by interpolation between two data points spanning 100sec-1 as follows        VIS100 = VIS1 + m (100-T1)                                                (3)

Where M =((VIS2-VIS1)/T2-T1))

VIS2 AND T2 are viscosity and shear rate at the first data point above 100 s-1 and VIS1 And T1 are viscosity and shear rate at the first data point below 100 s-1.

- Report

  - Report the rate index value as c, the viscosity as VIS100, correlation coefficient as r2, and intercept as ln b.  Label increasing and decreasing sweep data using INC and DEC.

  - Rate index  c =  dimensionless, 3 deciaml places, .xxx

  - VIS100 = mPa-sec, 2 decimal places xx.xx

  - r2, dimensionless, 4 decimal places, .xxxx

  - ln b = Pa, 4 decimal places, x.xxxx

15. Precision

15.1 Precision was found to be dependent on the mean value of the measured property.  The data in this section was derived from a valid 6 laboratory/6 sample interlaboratory study and a 3 laboratory/7 sample interlaboratory study which did not meet the ASTM guidelines due to the small number of labs participating.  All data was used to generate the repeatability and reproducibility data for rate index. The valid ILS data was used to generate the reproducibility data for VIS100 as including the 3 laboratory data skewed the reproducibility data significantly.

|  | r | R |
| --- | --- | --- |
| Rate index, inc | .158-.149x | .604-.588x |
| Rate index, dec | .037-.028x | 1.216-1.213x |
| Vis100, inc | -1.38+.12y | -1.10+.19y |
| Vis100, dec | -1.00+.11y | -.54+.20y |

Where x = mean value of rate index and y = mean value of vis100 in mPa-s

  - Sample Precision Calculations

The maximum value for rate index is 1.000.  This occurs for fluids which exhibit Newtonian behavior  as observed for fresh diesel oils within the shear rate range of measurement in this test method.  The rate index, inc  reproducibility for fresh oils would be :

.604 - .588 (1.000) = .016

and for a drain with a value of .900

.604 - .588(.900) = .075

The VIS100 precision values are valid  for viscosities within the scope of this method , 12-35 mPa-s.  A drain with a VIS100, inc value of 15.20 mPa-s would have a repeatability of

-1.38 +.12(15.20) = 0.44

and a reproducibility of

-1.10 +.19(15.20) = 1.79

ANNEX

(mandatory Information)

X1.

  - Typical  example to determine noise level and minimum shear rate value.

The data are plotted as shown and the minimum noise level determined.  Data below this level are deleted for analysis.  These examples are from two different cone and plate systems.

APPENDIX 1

(nonmandatory information)

X1.1Analysis Example of data table following directions of section 13.

| shear stress | shear rate | viscosity | ln(shear rate) | ln(shear stress) |
| --- | --- | --- | --- | --- |
| Pa | 1/s | Pa.s |  |  |
| 1.034 | 15.79 | 0.06551 | 2.759376828 | 0.033434776 |
| 1.673 | 31.58 | 0.05299 | 3.452524009 | 0.514618422 |
| 2.183 | 47.37 | 0.04608 | 3.857989117 | 0.780700078 |
| 2.634 | 63.16 | 0.04171 | 4.145671189 | 0.968503603 |
| 3.04 | 78.95 | 0.03851 | 4.368814741 | 1.111857515 |
| 3.436 | 94.74 | 0.03627 | 4.551136297 | 1.234308004 |
| 3.814 | 110.5 | 0.03451 | 4.705015521 | 1.338678507 |
| 4.166 | 126.3 | 0.03298 | 4.838660029 | 1.426956343 |
| 4.515 | 142.1 | 0.03177 | 4.956531035 | 1.507405187 |
| 4.859 | 157.9 | 0.03077 | 5.061961921 | 1.580832655 |
| 5.184 | 173.7 | 0.02985 | 5.157329673 | 1.645576959 |
| 5.51 | 189.5 | 0.02908 | 5.244389025 | 1.706564623 |
| 5.83 | 205.3 | 0.0284 | 5.324472324 | 1.763017 |
| 6.15 | 221.1 | 0.02782 | 5.398615088 | 1.816452082 |
| 6.478 | 236.8 | 0.02735 | 5.467215903 | 1.868411821 |
| 6.777 | 252.6 | 0.02683 | 5.53180721 | 1.913534526 |
| 7.08 | 268.4 | 0.02638 | 5.592478405 | 1.957273908 |
| 7.374 | 284.2 | 0.02595 | 5.649678216 | 1.9979603 |
| 7.676 | 300 | 0.02559 | 5.703782475 | 2.038098578 |

VIS100 = 36.27 + m(100-94.74) = 36.27 +M(5.26)

M = (34.51-36.27)/(110.5-94.74) = -.1117

VIS100 = 36.27 +(-.1117)(5.27)

=36.27 -.59

=35.68 mPa-s

rate index c = .677

r2= .9997

Intercept ln b = -1.8376 Pa   b = 0.1592

References

- “Rheology of Soot thickened Diesel Engine Oils”, K. Selby, SAE 981369, May 1998

- “Understanding SMOT through Designed Experimentation Part 3: An Improved approach to Drain Oil Viscosity Measurements – Rotational Rheology”, H.F.George, E.A. Bardasz and B. Soukup, SAE 97692,  May 1997